User Manual 
ROOT Logger						Version May 2022 (ahp1)

Caution:

Always use a 12 V battery. Depending on the type of the 12 V battery, it must be removed for transport; gel/liquid can leak if stored incorrectly.

The AGM battery currently installed should be leak-proof.

Before data collection:

Charge the battery (ON/OFF switch to OFF.)
Delete old files on the SD card.

Data Collection:

- Insert the SD card.

- Connect the sensors:  
Load Cell to LOAD socket (6),
Draw-wire sensor to CABLE socket (5)

- Set the ON/OFF switch (4) to ON.

(When switching on and off, a new .csv is generated in each case).

- Note the exact time of the measurement.

- BT-MON/LOG switch (1)

  - When the green LED (3) flashes, the measured values are written directly to the SD.

  - - When the blue LED flashes (2), the measurement values can be read out via Bluetooth module (HC-05-ii) or displayed on the laptop via USB cable using the USB socket (7).

  - If no LED flashes. See if the SD has been inserted correctly.

ATTENTION: The data is not written to the SD card. When the blue LED flashes.

Apps (only Android devices for now (not iPhone compatible))

App for Android:

Serial Bluetooth by Kai Morich

App for Windows

Serial port monitor (free version)– not yet tested

2

1

2

3

4

5

6

7

Read out the data

| Time elapsed since start [ms] | Unix Time Stamp | Cable [cm] | Load [N] | Battery voltage [V] |
| --- | --- | --- | --- | --- |
| 29239 | 1652178932 | 25 | -167.545 | 12.38 |
| 29249 | 1652178932 | 25 | -170.112 | 12.38 |
| … | … | … | … | … |

Unix time stamp to date/time

The time stamp is stored via the internal battery of the SD Shield.

In Excel:

- =(((A1/60)/60)/24)+DATE(1970;1;1)

- Format to: hh:mm:ss TT.MM.JJ (or whatever you prefer)

In Python:

import datetime

dt = datetime.datetime.fromtimestamp(1652178932)

# Don’t forget to set the time zone.

Technical data

At 12.2 volts, an average of 140 mA (1.16 mA idle), 
i.e., a fully charged battery with, for example, 6 Ah, lasts theoretically for 
approx. 40 hours (6 Ah/0.14A = 42.9 h).

Note that the lead-acid battery (AGM) is fully charged at a terminal voltage of 12.8 V (100 %) and discharged at a terminal voltage of approx. 11.8 V (0 %).

Schematic diagram

Calibration:

LEANE – DBBE 200 kg: factor 10050 – calibrated with a 20 kg and 40 kg load – May 2022

ME-Systeme - 5000 kg_ (factor 660 – measured with a 40 kg and 80 kg load – May 2022

Pin assignment

| LoadCell Leane DBBE (small – 200 kg) to HX711 |  |
| --- | --- |
| blue (Pin 1) | E+ |
| black (Pin 2) | E- |
| White (Pin 3) | A- |
| Red (Pin 4) | A+ |

| LoadCell Leane DBBE (large – 2000 kg) to HX711 |  |
| --- | --- |
| blue (Pin 1) | E+ |
| black (Pin 2) | E- |
| White (Pin 3) | A- |
| Red (Pin 4) | A+ |

For more information check: https://learn.sparkfun.com/tutorials/load-cell-amplifier-hx711-breakout-hookup-guide/all)

Parts List

| No° | Article | Description | Seller |
| --- | --- | --- | --- |
| 1 | Basetech Outdoor IP67 | Case 460 x 360 x 175 mm | conrad.ch |
| 2 | Neutrik NC5FD-LX-B | 5-Pol socket | conrad.ch |
| 3 | Neutrik NC5MX | 5- Pol plug | conrad.ch |
| 1 | Neutrik NAUSB-W-B | USB plug | conrad.ch |
| 1 | Neutrik SCDP-0CON | Seal for sockets | conrad.ch |
| 1 | TC-R13-208B-02 12 V/DC 20 A | Switch (green LED) | conrad.ch |
| 1 | TC-R13-208A-02 250 V/AC 10 A | Switch (no LED) | conrad.ch |
| 1 | Arduino Uno | Microcontroller | conrad.ch |
| 1 | TC-9927152 | HC-05 Bluetooth module (Android only) | conrad.ch |
| 1 | Joy-it SEN-HX711-20 | HX 711 - Amplifier board for the load cell | conrad.ch |
| 1 | Adafruit Assembled Data Logging shield | SD Shield Logger for Arduino Uno | conrad.ch |
| 1 | SD Card | 32 GB | conrad.ch |
| 1 | Motobatterie YTX7A-BS Okay | Battery 12 V / 6Ah | Landi.ch |
| 1 | Kemo Spannungswandler (3 - 15 V/DC 1.5 A) | Voltage transformer/stabilizer | conrad.ch |
| 1 | 110x80x70mm IP67 | Transparent plastic housing Arduino | bastelgarage.ch |